HIC.6.c / HIC.6.d (stop-work authority) — No instrument is used from a load with a failed biological indicator, a failed Bowie-Dick test, or a steriliser alarm mid-cycle; no single-use device is reprocessed outside the approved list or beyond its maximum cycles.

HIC.6 Record 06 — Instrument stop-work and refusal log

What is this record, in plain terms? This record is the proof a suspect instrument was actually refused on the spot by whoever noticed it -- a CSSD technician, an OT nurse, a clinician -- rather than used because the case was already under way.

A real example of when you'd use this An OT nurse notices a pack's chemical indicator did not change colour and refuses to open it on the operating table, calling for a replacement set from a verified load instead.

Why does NABH ask for this? Using an instrument "just this once" while the recall investigation is still under way defeats the entire purpose of running the investigation in the first place -- the stop has to hold until the load is cleared, not until it is convenient.

How to fill this in — column by column

Date/Time: When identified.

Staff Member / Role: Who refused it.

Item / Issue Identified: What was wrong.

Replacement Used / Outcome: What happened next.

The first data row below (shaded yellow, and marked WORKED EXAMPLE in the row itself) is a worked example already filled in correctly. Delete it or overwrite it once you start using this record for real — don't leave the example row mixed in with real entries. Add more blank rows as needed; don't create a second copy of this file.

Where this requirement comes from (verified, clickable sources)

• National Accreditation Board for Hospitals and Healthcare Providers. (2022). Standards for Small Healthcare Organisations (3rd ed.). NABH SHCO Accreditation Programme page. Hospital Infection Control chapter, standard HIC.6.

| Date/Time | Staff Member / Role | Item / Issue Identified | Replacement Used / Outcome |
| --- | --- | --- | --- |
| WORKED EXAMPLE — 18-Mar-2026, 9:40 AM | Nurse T. George (OT) | Chemical indicator did not change colour | Replacement set opened from a verified load; suspect set quarantined |